Heart Rate (BPM) Analysis Engine

Engineering Specification - High-Priority Analysis Sets & Mathematical Definitions

.NET MAUI | FastAPI | Resting/Spot Measurements First | Device-Ready Architecture

| Document | Version | Date | Status |
| --- | --- | --- | --- |
| Engineering Specification | 0.2 Draft | September 13, 2026 | For implementation review |

Core engineering decision Heart rate will use the shared Vitals statistics/visualization infrastructure, but interpretation will be heart-rate-specific and context-aware. BPM measures rate, not rhythm: Vitals may summarize longitudinal patterns, symptom timing, medication changes, and cross-vital context, but it must not diagnose arrhythmias from spot BPM values.

This specification is an implementation companion to the Vitals Phase 4 Clinical Visualization Engine. It refines the roadmap concept that a common mathematical engine can serve multiple vital signs by defining heart-rate-specific context, eligibility rules, formulas, safety language, and result contracts. The primary launch target is manually entered or device-reported spot heart-rate measurements obtained with the daily Vitals reading workflow. [1]

1. Purpose and Scope

This document defines the high-priority heart-rate analysis sets for Vitals, including the mathematical formulas used to calculate each value, the minimum data required to activate each analysis, the proposed data model, and the presentation guardrails for the mobile UI and physician-grade PDF.

The initial implementation is designed for spot measurements such as a pulse value from a blood-pressure monitor, pulse oximeter, or manual entry. Dense wearable data may be supported later, but P0 analysis must remain useful without smartwatch or ECG integrations.

Not a rhythm diagnostic system A BPM value describes heart rate only. Tachyarrhythmias and bradyarrhythmias are rhythm diagnoses that generally require ECG/rhythm data and clinical interpretation. Vitals may report observed rate patterns and associations; it must not infer atrial fibrillation, SVT, heart block, sinus-node dysfunction, or other rhythm diagnoses from BPM alone.

Citation convention Bracketed citations such as [2] refer to the numbered References block in Section 17. Clinical thresholds, population reference ranges, diagnostic/safety boundaries, and externally validated formulas carry adjacent citations. Sample-count gates, time windows, display thresholds, grouping rules, and UI activation criteria that are not externally sourced are explicitly treated as Vitals engineering decisions.

2. Clinical and Engineering Guardrails

• Resting context is mandatory for threshold and longitudinal resting-rate interpretation. Activity, posture, symptoms, fever, stress, medications, and conditioning can materially alter heart rate.[2][3]

• For most adults, a resting heart rate of 60-100 BPM is commonly described as typical. A resting rate above 100 BPM is commonly termed tachycardia; a rate below 60 BPM is commonly termed bradycardia, but lower rates may be physiologic in athletes, during sleep, or with medications.[2][3][4]

• The 2018 ACC/AHA/HRS bradycardia guideline notes that population studies often use <50 BPM in sinus-node dysfunction definitions, but there is no single minimum heart rate at which treatment is automatically indicated. Symptom correlation is important.[4]

• Vitals should describe observations, not assign diagnoses. Preferred wording: "4 resting readings above 100 BPM were recorded" rather than "you have tachycardia."

• Spot-reading dispersion is not Heart Rate Variability (HRV). True HRV requires beat-to-beat RR/NN interval data. The spot-reading module may report SD/IQR between resting measurements, but must not label those values HRV. [9]

• Manual spot measurements support "percentage of logged readings" metrics, not duration-based burden. Time above/below a rate threshold requires dense continuous or near-continuous data.

• Medication and symptom associations must be expressed temporally ("occurred after" / "occurred near") rather than causally ("was caused by").

• Published thresholds are reference defaults only. Patient- or clinician-defined resting targets must override display defaults when configured.

3. Priority Summary

P0 analysis sets form the launch package for manual or device-reported spot readings. P1 adds a structured orthostatic measurement workflow. P2 items are intentionally deferred until Vitals receives dense wearable or beat-to-beat data.

| Priority | Analysis set | Spot / manual | Dense / structured source |
| --- | --- | --- | --- |
| P0 | Latest rate + context | Yes | Yes |
| P0 | Resting central tendency | Yes | Yes |
| P0 | Resting trend: OLS + LOESS | Yes | Yes |
| P0 | Personal baseline deviation | Yes | Yes |
| P0 | High / low resting-rate events | Yes | Yes |
| P0 | Between-reading dispersion | Yes | Yes |
| P0 | Time-of-day patterns | Yes | Yes |
| P0 | Symptom association | Yes | Yes |
| P0 | Medication-change association | Yes | Yes |
| P0 | Cross-vital same-event context | Yes | Yes |
| P0 | Data-support / density indicators | Yes | Yes |
| P1 | Orthostatic HR + BP response | Guided sequence | Guided/device sequence |
| P2 | Duration-based rate burden | No | Dense data only |
| P2 | True HRV (SDNN/RMSSD) | No | RR/NN intervals only |
| P2 | Heart-rate recovery | Structured exercise only | Structured exercise only |

4. Required Heart-Rate Data Model

Heart rate is only analytically meaningful when Vitals knows enough about the measurement context to decide whether observations are comparable. Existing records can migrate safely with nullable context fields; unknown-context values remain visible in History but should be excluded from context-specific claims until the user or source provides sufficient metadata.

| Field | Type / example | Required | Purpose |
| --- | --- | --- | --- |
| bpm | integer; 74 | Yes | Primary heart-rate value. |
| occurred_at_utc | timestamp | Yes | Canonical ordering/windowing. |
| local_offset_minutes | integer | Yes | Correct local time-of-day bucketing. |
| activity_context | resting / post_activity / during_activity / unknown | Recommended | Defines eligibility for resting analysis. |
| posture | seated / supine / standing / unknown | Recommended | Required for orthostatic and posture-aware comparison. |
| symptom_tags | dizziness, palpitations, syncope, dyspnea, fatigue, chest discomfort, other | Optional | Enables symptom association. |
| source_type | manual / BP_monitor / pulse_oximeter / wearable / imported | Yes | Provenance and capability gating. |
| measurement_event_id | UUID | Recommended | Links HR with BP, SpO2, temperature, weight captured together. |
| device_irregular_pulse_flag | bool? | Optional | Preserve device-supplied flag; do not infer from BPM. |
| notes | text | Optional | Clinician/patient context. |
| is_invalidated | bool | Yes | Exclude known-bad readings without destructive deletion. |

4.1 Comparable resting set

Default P0 resting analyses should use observations where activity_context = resting and is_invalidated = false. Standing observations should be excluded from ordinary resting trend unless explicitly requested; they belong to posture/orthostatic analysis. Unknown-context readings may be shown as raw history but should not drive resting threshold or baseline claims.

If the source cannot distinguish rest from activity, Vitals should ask the user to confirm context during entry rather than silently assuming every BPM value is resting.

4.2 Measurement-event linkage

A single daily Vitals entry should generate or reuse one measurement_event_id for BP, heart rate, SpO2, temperature, and weight. Cross-vital analysis should prefer exact event linkage. A nearest-timestamp fallback is allowed only when values are within 2 minutes and belong to the same patient.

5. Mathematical Conventions

Unless otherwise stated, y_i is heart rate in BPM and x_i is elapsed time in days from the start of the selected analysis window. Calculations use double precision internally. No missing heart-rate values are imputed for spot-reading analysis.

| Windowing | Use [start, end] inclusive by timestamp. The UI may offer 7-, 14-, 30-, and 60-day windows. |
| --- | --- |
| Current value | Latest non-invalidated observation by occurred_at_utc. |
| Display rounding | Current BPM: integer. Means/medians: 1 decimal in clinician view, nearest integer on compact cards. Slopes: 0.1 BPM/day. Percentages: 1 decimal. |
| Duplicate handling | Do not average independent readings automatically. If a device/source explicitly identifies repeats as one measurement session, retain all raw values and optionally expose a source-derived final value. |
| Distinct-day count | Count unique local calendar dates represented by eligible readings. |
| No interpolation | Spot readings are observations, not a continuous waveform. Never infer unobserved duration above/below a threshold from sparse data. |

6. P0 Analysis Sets - Spot / Resting Measurements

The following analysis sets should be implemented first. Proposed sample gates are engineering defaults for product stability and are not clinical diagnostic criteria.

6.1. Snapshot and Measurement Context

P0 - Ship first   Spot + dense sources

| Purpose | Provide the most recent heart-rate observation with enough context for a physician to interpret what the number represents. |
| --- | --- |
| Required inputs | Latest valid heart-rate observation; activity context; posture; symptoms; source; timestamp; linked measurement event. |
| Calculation / formulas | latest = argmax_t(reading_i) over valid observations. No statistical transformation is applied to the latest value. |
| Activation gate | One valid observation. |
| Analysis output | BPM, timestamp, resting/activity context, posture, source, symptom tags, linked BP/SpO2/temperature when present. |
| Physician-facing example | "74 BPM, resting/seated, Sep 12 at 7:14 AM." |
| Engineering notes | Do not attach a diagnostic label to a single rate. If context is unknown, say so explicitly. |

6.2. Resting Central Tendency and Range

P0 - Ship first   Comparable resting readings

| Purpose | Summarize the patient's observed resting-rate level over a selected window while preserving the distribution. |
| --- | --- |
| Required inputs | Eligible resting BPM values y_1...y_n in the selected window. |
| Calculation / formulas | Arithmetic mean: y_bar = (1/n) * Σ y_i Median: sort(y); odd n -> middle value; even n -> average of two middle values. Minimum: min(y_i) Maximum: max(y_i) Range: max(y_i) - min(y_i) [8] |
| Activation gate | Descriptive summary: n >= 3 eligible readings across >= 3 distinct local dates. Latest value remains available below this gate. |
| Analysis output | n, distinct days, mean, median, minimum, maximum, range for 7/14/30/60-day windows. |
| Physician-facing example | "30-day resting average 72.4 BPM; median 71 BPM; observed range 58-94 BPM." |
| Engineering notes | Mean is familiar; median is more resistant to isolated extremes and is preferred for personal-baseline comparisons. |

6.3. Between-Reading Dispersion

P0 - Ship first   Comparable resting readings

| Purpose | Quantify how spread out spot resting-rate observations are without mislabeling the result as beat-to-beat HRV. |
| --- | --- |
| Required inputs | Eligible resting BPM values in the selected window. |
| Calculation / formulas | Sample SD: s = sqrt[ Σ(y_i - y_bar)^2 / (n - 1) ] IQR = Q3 - Q1, using linear percentile interpolation: h=(n-1)p; Q(p)=y_floor(h)+(h-floor(h))*(y_ceil(h)-y_floor(h)). Optional descriptive CV: CV_spot = 100 * s / y_bar. [8] |
| Activation gate | n >= 5 for SD/IQR display. CV_spot may be retained internally but should not be labeled HRV. |
| Analysis output | Sample SD, IQR, range, optional CV_spot, plus comparison with the prior equal-length period. |
| Physician-facing example | "Resting readings were more spread out this month: SD 8.2 BPM vs 5.1 BPM in the prior month." |
| Engineering notes | Avoid universal "good/bad variability" thresholds for sparse spot readings. Show the statistic and longitudinal change rather than diagnosing autonomic/rhythm abnormalities. |

6.4. Resting Trend: Ordinary Least Squares

P0 - Ship first   Comparable resting readings

| Purpose | Measure the direction and magnitude of longitudinal change in resting heart rate using a reproducible linear model. |
| --- | --- |
| Required inputs | Eligible resting readings (x_i days, y_i BPM) in one analysis window. |
| Calculation / formulas | Slope: b1 = Σ[(x_i-x_bar)(y_i-y_bar)] / Σ[(x_i-x_bar)^2] Intercept: b0 = y_bar - b1*x_bar Predicted: y_hat_i = b0 + b1*x_i SSE = Σ(y_i-y_hat_i)^2; SST = Σ(y_i-y_bar)^2 R^2 = 1 - SSE/SST (undefined when SST=0) MSE = SSE/(n-2) SE(b1) = sqrt[MSE / Σ(x_i-x_bar)^2] t = b1 / SE(b1); df=n-2; p = 2*[1 - F_t(¦t¦; df)] Modeled window change = b1 * (x_max - x_min). [6] |
| Activation gate | Directional slope: n >= 5 across >= 5 distinct days and span >= 7 days. Display p-value only when n >= 8 and span >= 14 days. Product gates are proposed engineering defaults. |
| Analysis output | Slope in BPM/day, modeled change over observed span, R^2, optional p-value/CI for clinician/PDF view. |
| Physician-facing example | "Resting heart rate increased by an estimated 7.8 BPM over the observed 30-day span (slope +0.27 BPM/day)." |
| Engineering notes | Statistical significance is not clinical significance. Consumer UI should emphasize direction/magnitude; R^2 and p-value belong in clinician details/PDF. [6] |

6.5. LOESS Smoothed Trend

P0 - Ship first   Visualization only

| Purpose | Provide a smooth visual trajectory without forcing a straight-line relationship, while retaining every raw observation. |
| --- | --- |
| Required inputs | Eligible resting readings (x_i, y_i). |
| Calculation / formulas | Use local linear LOESS with span q=0.60. For each evaluation x0, select k=max(5, ceil(0.60*n)) nearest points. Let d be the largest ¦x_i-x0¦ among selected points and u_i=¦x_i-x0¦/d. Tricube weight: w_i=(1-¦u_i¦^3)^3 for ¦u_i¦<1; otherwise 0. Fit weighted local line y=a+b*x by minimizing Σ w_i[y_i-(a+b*x_i)]^2; evaluate at x0. [7] |
| Activation gate | n >= 7 eligible observations and at least 7 days of span. If d=0 or geometry is singular, omit smoothing at that point. |
| Analysis output | Smoothed series for chart overlay; raw points always remain visible. |
| Physician-facing example | Chart shows the raw resting BPM measurements with a smoothed trajectory. |
| Engineering notes | Do not use robust iterations that automatically suppress physiologic extremes in the first implementation; an extreme reading may be clinically important. LOESS is visual support, not a diagnosis. |

6.6. Personal Baseline Deviation

P0 - Ship first   Comparable resting readings

| Purpose | Highlight change relative to the patient's established resting-rate baseline even when recent values remain inside a broad population reference range. |
| --- | --- |
| Required inputs | Baseline resting set and recent resting set. |
| Calculation / formulas | Default recent window R = last 7 days. Default baseline B = the immediately preceding 30 days (days -37 through -8 relative to now). Baseline median M_B = median(B); recent median M_R = median(R). Absolute shift: Delta_BPM = M_R - M_B Percent shift: Delta_% = 100*(M_R-M_B)/M_B, when M_B != 0. Optional standardized shift: z = (M_R - mean(B))/s_B when baseline n>=7 and s_B>0. |
| Activation gate | Baseline >= 7 resting readings across >= 7 distinct days; recent >= 3 resting readings across >= 3 distinct days. If insufficient, suppress baseline claim. |
| Analysis output | Baseline median, recent median, absolute BPM change, percent change, optional z-score internally. |
| Physician-facing example | "Recent resting heart rate is 12 BPM higher than the preceding 30-day baseline." |
| Engineering notes | Do not label a baseline shift dangerous/abnormal solely from Delta_BPM. Physician context, symptoms, temperature, hydration, activity, and medication changes matter. |

6.7. High / Low Resting-Rate Event Analysis

P0 - Ship first   Resting readings only

| Purpose | Count repeated observed rates outside reference or clinician-defined targets without converting sparse readings into duration-based burden. |
| --- | --- |
| Required inputs | Eligible resting BPM values; target profile; timestamps. |
| Calculation / formulas | Default descriptive reference thresholds: high if y_i > 100 BPM; low if y_i < 60 BPM; marked-low reference if y_i < 50 BPM. [2][3][4] Count: N_A = Σ I(condition_i) Logged-reading proportion: P_A = 100 * N_A / n Episode clustering (optional): sort qualifying readings; readings of the same threshold direction separated by <=15 minutes belong to one episode. New episode when gap >15 minutes. |
| Activation gate | Counts: n >= 1. Percentages: n >= 5 resting readings in selected window. Use clinician/patient custom thresholds when configured. |
| Analysis output | High/low reading count, percentage of logged resting readings, episode count, dates/times, extrema, symptom/context links. |
| Physician-facing example | "4 of 28 logged resting readings were above 100 BPM; 2 occurred near reported palpitations." |
| Engineering notes | Never convert this to "time above 100" for spot data. AHA notes that low rates may be normal in athletes/sleep/medication contexts, and the bradycardia guideline stresses that symptoms and clinical context matter. [2][4] |

6.8. Time-of-Day Pattern Analysis

P0 - Ship first   Comparable resting readings

| Purpose | Reveal repeatable morning/afternoon/evening/overnight differences that may help a physician interpret medication timing, symptoms, sleep, and daily physiology. |
| --- | --- |
| Required inputs | Eligible resting readings with reliable local time. |
| Calculation / formulas | Default local-time buckets: Morning 05:00-11:59; Afternoon 12:00-16:59; Evening 17:00-21:59; Overnight 22:00-04:59. For each bucket b: mean_b = (1/n_b)Σy_i; median_b = median(y_i). Difference from overall median: Delta_b = median_b - median_all. Highest qualified period = argmax_b(mean_b). |
| Activation gate | Per bucket: >=3 readings across >=3 distinct days. Pattern summary requires at least 2 qualified buckets. |
| Analysis output | Bucket n, mean, median, range, delta from overall median; highest/lowest qualified period. |
| Physician-facing example | "Evening resting readings averaged 8 BPM higher than morning readings during the last 30 days." |
| Engineering notes | Bucket cutoffs are product defaults, not clinical standards. Preserve local-time zone/offset from the original observation. |

6.9. Symptom Association

P0 - Ship first   Spot + symptom events

| Purpose | Help physicians establish temporal relationships between symptoms and heart-rate observations, an important principle in bradycardia evaluation. [4] |
| --- | --- |
| Required inputs | Heart-rate observations, symptom tags/events, timestamps, threshold profile. |
| Calculation / formulas | Direct association when symptom_tags are stored on the same measurement event. Otherwise, for a standalone symptom timestamp s_j, choose nearest eligible HR reading h_i if ¦t_i-s_j¦ <= 15 minutes. Associated count for symptom k: A_k = Σ I(symptom_k has linked HR). Low-associated proportion: P_low,k = 100 * N(symptom_k linked to HR below low threshold) / A_k. High-associated proportion defined analogously. |
| Activation gate | Individual associations may display with one event. Aggregate statement requires >=2 associated symptom events of the same category. The ±15-minute fallback is a product default, not a physiologic causality window. |
| Analysis output | Event list, associated BPM, time difference, posture/activity, high/low status, aggregate counts by symptom. |
| Physician-facing example | "2 of 3 dizziness episodes with nearby heart-rate data occurred with a resting rate below 60 BPM." |
| Engineering notes | Use "occurred near" / "was recorded with," never "caused." The ACC/AHA/HRS guideline emphasizes symptom-rate correlation but also notes that symptoms can have competing etiologies. [4] |

6.10. Medication-Change Association

P0 - Ship first   Medication history + resting readings

| Purpose | Show whether resting-rate measurements changed after a documented medication start, stop, or dose/frequency change, without asserting drug causation. |
| --- | --- |
| Required inputs | Medication change timestamp; eligible resting readings; optional medication class; other medication-change events. |
| Calculation / formulas | Default windows: PRE = [change-14d, change); POST = (change, change+14d]. M_pre = median(PRE); M_post = median(POST). Absolute difference: Delta_med = M_post - M_pre Percent difference: Delta_med_% = 100*(M_post-M_pre)/M_pre. Optional mean difference uses the same equation with arithmetic means. |
| Activation gate | >=3 resting readings across >=3 distinct days in both PRE and POST. Mark comparison confounded if another medication change occurs anywhere in the combined 28-day comparison window. |
| Analysis output | Pre/post n, median/mean BPM, absolute/percent difference, medication event details, confounded flag. |
| Physician-facing example | "Median resting heart rate was 10 BPM lower in the 14 days after the recorded dose change than in the prior 14 days." |
| Engineering notes | The bradycardia guideline specifically recommends careful medication review because drugs can elicit or worsen slow rates. Present association only; physician determines causality. [4] |

6.11. Cross-Vital Same-Event Context

P0 - Ship first   HR + other Vitals measurements

| Purpose | Place heart-rate observations beside BP, temperature, SpO2, and other readings captured at the same time so physicians can interpret physiologic context. |
| --- | --- |
| Required inputs | measurement_event_id or timestamp-matched observations; heart-rate exception predicate; other-vital exception predicates supplied by each metric-specific strategy. |
| Calculation / formulas | Primary co-occurrence count for HR condition H and other-vital condition O: C_HO = Σ I(H_i AND O_i) over aligned measurement events. Co-occurrence percentage: P_HO = 100*C_HO / N_H, where N_H is the number of HR-condition events that have the paired vital available. Optional physician-only Pearson correlation for paired continuous values (n>=10): r = Σ[(x_i-x_bar)(y_i-y_bar)] / sqrt[Σ(x_i-x_bar)^2 * Σ(y_i-y_bar)^2]. |
| Activation gate | Co-occurrence: >=1 linked event. Percentage: >=5 HR-condition events with paired vital. Correlation: >=10 paired events across >=7 days; hide if either variable has zero variance. |
| Analysis output | Event table with HR + BP + SpO2 + temperature; co-occurrence counts; optional r in clinician detail. |
| Physician-facing example | "Three elevated resting-rate readings occurred in measurement events that also contained an elevated temperature." |
| Engineering notes | Do not hardcode temperature/SpO2/BP clinical thresholds inside HeartRateAnalysisStrategy. Query those metric strategies for their own exception states. Correlation does not imply causation. |

6.12. Data Support and Density

P0 - Ship first   All spot datasets

| Purpose | Make it obvious when Vitals has enough observations for a particular claim and prevent sparse data from producing authoritative-looking statistics. |
| --- | --- |
| Required inputs | Eligible reading count, unique local dates, observed span, selected window length. |
| Calculation / formulas | Distinct-day coverage: D = 100 * (# distinct local dates with >=1 eligible reading) / (# calendar days in selected window). Observed span: S = (max timestamp - min timestamp) in days. |
| Activation gate | Support states: Snapshot = n>=1. Descriptive = n>=3 across >=3 days. Trend = n>=5 across >=5 days and S>=7. LOESS = n>=7. Significance display = n>=8 and S>=14. Baseline shift uses its own gate. |
| Analysis output | Support state, n, distinct days, span, distinct-day coverage, list of unavailable analyses/reasons. |
| Physician-facing example | "30-day summary uses 18 resting readings collected on 16 days." |
| Engineering notes | Call this data support/coverage, not confidence in a diagnosis. Manual coverage is not the same as continuous monitoring coverage. |

7. P1 Structured Analysis - Orthostatic HR + BP Response

Orthostatic analysis is valuable because Vitals already captures both heart rate and blood pressure. It should be implemented as a guided measurement sequence rather than inferred from unrelated standing and seated readings. A practical protocol measures BP and heart rate after approximately 5 minutes supine and again after standing; many workflows include measurements at 1 and 3 minutes. [5]

Scope boundary Vitals may calculate and display orthostatic changes. The first implementation should not diagnose POTS, orthostatic hypotension, autonomic failure, or another cause. AHA defines orthostatic hypotension by a sustained BP reduction of at least 20 mmHg systolic or 10 mmHg diastolic within 3 minutes of standing; that reference can be shown to clinicians, but diagnostic wording should undergo clinical review before any patient-facing label is shipped. [5]

7.1. Orthostatic Response

P1 - Structured / gated   Guided supine-to-standing sequence

| Purpose | Give physicians objective paired changes in heart rate and blood pressure during a controlled posture transition. |
| --- | --- |
| Required inputs | Supine/rested HR + BP; standing HR + BP at 1 minute and 3 minutes; exact timestamps; symptom tags. |
| Calculation / formulas | For time point t in {1m,3m}: Delta_HR_t = HR_standing,t - HR_supine Delta_SBP_t = SBP_standing,t - SBP_supine Delta_DBP_t = DBP_standing,t - DBP_supine Percent HR change_t = 100*Delta_HR_t/HR_supine. Optional compensation ratio for research/clinician review only: R_t = Delta_HR_t / abs(Delta_SBP_t), undefined when Delta_SBP_t=0. |
| Activation gate | All paired values required for a complete time point. Flag timing deviation if standing measurement differs from target by >30 seconds. Require 5-minute supine rest confirmation in the guided UI. |
| Analysis output | Supine/1m/3m values, each delta, symptoms during transition, timing-quality flag. |
| Physician-facing example | "Supine 68 BPM -> standing 88 BPM at 1 minute (+20); SBP 128 -> 112 mmHg (-16)." |
| Engineering notes | AHA notes HR normally increases as part of compensation on standing and distinguishes orthostatic hypotension from postural tachycardia syndromes. Do not infer etiology from these deltas alone. [5] |

8. P2 / Deferred Device-Enabled Analysis

The following analyses should be represented in capability contracts now but not activated from ordinary spot BPM values.

| Deferred analysis | Required source | Mathematical / engineering note |
| --- | --- | --- |
| Duration-based high/low rate burden | Dense wearable/telemetry sampling | time_above = 100 * observed_duration_above / valid_observed_duration. Never derive from sparse spot readings. |
| True Heart Rate Variability - SDNN | Validated RR/NN interval series | SDNN = sqrt[Σ(NN_i-NN_bar)^2/(N-1)]. Do not compute from daily BPM values. [9] |
| True Heart Rate Variability - RMSSD | Validated RR/NN interval series | RMSSD = sqrt[Σ(NN_{i+1}-NN_i)^2/(N-1)]. [9] |
| Heart-rate recovery | Structured exercise protocol with peak/end HR and timed recovery HR | HRR_1 = HR_peak_or_end - HR_1min. Protocol details must be standardized before interpretation. [10] |
| Rhythm classification | ECG or validated rhythm-capable device output | Out of scope for BPM mathematics; preserve device/ECG result provenance and do not infer rhythm from rate alone. |

9. Result Contract and Service Architecture

Heart-rate analysis should be implemented as a metric-specific strategy over a shared statistical toolkit. The strategy owns clinical grouping and eligibility; shared services own reusable math such as mean/median, OLS, LOESS, SD, IQR, and correlation.

Proposed result DTO HeartRateAnalysisResult period latest resting_summary { n, distinct_days, mean, median, min, max, range } dispersion { sd, iqr, spot_cv? } trend { slope_bpm_per_day, modeled_change, r2, p_value?, loess[] } baseline_shift { baseline_median, recent_median, delta_bpm, delta_percent } threshold_events { high, low, marked_low, custom_target_exceptions } time_of_day[] symptom_associations[] medication_associations[] cross_vital_context[] orthostatic? data_support { state, n, distinct_days, span_days, day_coverage_percent } disclosures[]

• IHeartRateAnalysisStrategy is responsible for filtering comparable resting observations and applying heart-rate-specific threshold/wording rules.

• IStatisticsService provides deterministic implementations of mean, median, quantiles, sample SD, OLS, Student-t p-value, LOESS, and correlation.

• IMedicationTimelineService supplies normalized start/stop/dose/frequency events and overlap/confounding detection.

• ISymptomAssociationService resolves same-event or nearest-timestamp associations without embedding diagnostic logic.

• ICrossVitalContextService aligns measurements by measurement_event_id first, then by the strict timestamp fallback.

• UI and PDF layers consume the same result model so displayed numbers cannot drift between app and exported report.

10. Recommended UI and Physician PDF Composition

10.1 Dashboard card

Heart Rate Latest: 74 BPM - Resting - Today 7:14 AM 30-day resting average: 72 BPM Personal baseline: +5 BPM vs preceding 30 days Trend: Gradually increasing

The Dashboard should show only a compact subset: latest value/context, one longitudinal summary, and one meaningful deviation/trend cue. Detailed event counts, symptoms, medications, and statistical outputs belong on the Heart Rate Details screen and physician PDF.

10.2 Heart Rate Details

• Latest resting rate + measurement context.

• 7/14/30/60-day central tendency and observed range.

• Raw resting points + LOESS overlay; OLS direction/magnitude in details.

• Personal-baseline shift.

• High/low resting-rate reading counts with dates and symptoms.

• Time-of-day pattern cards only when data gates are met.

• Medication-change comparisons, with confounded comparisons clearly labeled.

• Linked BP/temperature/SpO2 context for selected outlier events.

• Orthostatic response module only after a guided measurement sequence exists.

10.3 Physician PDF

Example physician summary HEART RATE - 30 DAY SUMMARY Resting average: 72.4 BPM ¦ Median: 71 BPM ¦ Range: 48-108 BPM Trend: +0.27 BPM/day; modeled +7.8 BPM across observed span Personal baseline: recent median 77 BPM vs prior baseline 65 BPM (+12 BPM) Resting readings >100 BPM: 3 of 28 ¦ <60 BPM: 5 of 28 ¦ <50 BPM: 1 of 28 Symptoms: 2 low-rate readings occurred near reported dizziness Medication timeline: median resting HR 10 BPM lower after recorded dose change (temporal association; causality not inferred).

11. Language and Safety Rules

| Preferred wording | Do not say |
| --- | --- |
| "Four resting readings above 100 BPM were recorded." | "You have tachycardia." |
| "Recent resting rate is 12 BPM higher than the preceding baseline." | "Your heart condition is worsening." |
| "The lower readings occurred after the medication change." | "The medication caused bradycardia." |
| "Two dizziness episodes occurred near low heart-rate readings." | "Your low heart rate caused dizziness." |
| "Between-reading SD was 8.2 BPM." | "HRV was 8.2 BPM" from spot readings. |
| "The rate was recorded with an elevated temperature." | "The fever caused the high heart rate." |

Escalation/urgent-care messaging is intentionally outside this engineering specification because it becomes a clinical triage feature. Any future symptom/rate safety alerts require separate clinical review, product policy, and regulatory assessment.

12. Deterministic Test Vectors

The backend and any client-side mirror implementation should pass the same reference vectors. Values below are intentionally simple so failures are obvious.

| Vector | Input | Expected output |
| --- | --- | --- |
| A - central tendency / OLS | Days x=[0,1,2,3,4]; BPM y=[68,70,72,74,76] | mean=72; median=72; min=68; max=76; range=8; sample SD=sqrt(10)=3.1623; slope=2.0 BPM/day; intercept=68; R^2=1.0; modeled change=8 BPM. |
| B - threshold counts | Resting BPM [58,59,62,104,108,72] | n=6; low<60: 2 (33.3%); high>100: 2 (33.3%); marked-low<50: 0. |
| C - personal baseline | Baseline [64,66,65,67,63]; recent [76,78,77] | baseline median=65; recent median=77; Delta=+12 BPM; Delta%=18.4615%. |
| D - medication association | PRE [78,80,79,77]; POST [68,70,69,67] | M_pre=78.5; M_post=68.5; Delta=-10.0 BPM; Delta%=-12.7389%. |
| E - symptom window | Dizziness 10:00; HR 52 at 10:04. Palpitations 15:00; HR 112 at 15:12. HR 118 at 15:20. | With +/-15 min fallback: 52 and 112 are associated; 118 at +20 min is not selected for the 15:00 symptom. |
| F - orthostatic deltas | Supine HR 68, BP 128/76; 1m HR 88, BP 112/68; 3m HR 82, BP 116/70 | 1m: DeltaHR=+20, DeltaSBP=-16, DeltaDBP=-8. 3m: DeltaHR=+14, DeltaSBP=-12, DeltaDBP=-6. |

13. Verification Requirements

• Unit-test every formula and every eligibility gate independently.

• Run date/time tests around midnight, daylight-saving transitions, and time-zone changes; time-of-day buckets must use the reading's original local offset.

• Verify invalidated readings are excluded from calculations but remain auditable in History.

• Verify unknown-context readings never enter resting trend, resting threshold, or baseline calculations by default.

• Verify threshold boundary semantics exactly: default high is >100; low is <60; marked-low reference is <50. A reading of exactly 100 is not in the default high set; exactly 60 is not in the low set.

• Verify custom clinician/patient target thresholds replace reference thresholds without changing stored raw readings.

• Verify OLS against a trusted numerical library and compare R^2, slope SE, t statistic, and p-value within tight tolerance.

• Verify LOESS weighting/neighbor selection against a reference implementation using the same span and local-linear degree.

• Verify same-event cross-vital linkage takes precedence over timestamp matching.

• Verify medication comparisons are marked confounded when another medication change overlaps the comparison interval.

• Verify UI and PDF consume the same analysis DTO and produce the same displayed values after rounding.

14. Recommended Implementation Order

| Step | Engineering work | Outcome |
| --- | --- | --- |
| 1 | Extend heart-rate observation model | Add context/posture/source/symptoms/event linkage while preserving existing BPM records. |
| 2 | Extract shared statistics library | Deterministic mean/median/quantiles/SD/OLS/t/LOESS/correlation with reference tests. |
| 3 | Build resting summary + data support | Snapshot, 7/14/30/60 summaries, range, SD/IQR, support gates. |
| 4 | Build trend + personal baseline | OLS, LOESS, baseline comparison, safe language. |
| 5 | Build event + time-of-day analysis | High/low reading counts, custom targets, local-time buckets. |
| 6 | Add symptom + medication association | Timestamp/event linkage and pre/post comparison with confounding flags. |
| 7 | Add cross-vital context | Link HR to BP/SpO2/temperature using measurement_event_id. |
| 8 | Update Dashboard / Details / PDF | Metric-driven composition; same result DTO everywhere. |
| 9 | Add guided orthostatic workflow | P1 paired HR+BP measurements after P0 is stable. |
| 10 | Land future capability flags | Reserve dense-rate, RR/NN HRV, exercise recovery, and rhythm-source contracts without exposing unsupported metrics. |

15. Engineering Decisions to Lock

| Decision | Recommended value | Status |
| --- | --- | --- |
| Resting context required for trend/thresholds | Yes | Adopt now |
| Default resting reference high | >100 BPM | Reference only; clinically review wording |
| Default resting reference low | <60 BPM | Reference only; clinically review wording |
| Marked-low physician reference | <50 BPM | Reference only; no diagnosis |
| OLS trend minimum | 5 readings / 5 days / >=7-day span | Proposed engineering gate |
| Regression significance display | 8 readings / >=14-day span | Proposed engineering gate |
| LOESS span | 0.60, local linear, no robust outlier suppression | Proposed; validate visually |
| Personal recent window | 7 days | Proposed |
| Personal baseline window | preceding 30 days | Proposed |
| Symptom timestamp fallback | nearest HR within +/-15 minutes | Proposed product rule |
| Medication comparison window | 14 days before vs 14 days after | Proposed product rule |
| Cross-vital fallback alignment | same event first; otherwise <=2 minutes | Proposed product rule |
| Spot dispersion label | SD/IQR / "between-reading spread" | Adopt; never call HRV |
| Orthostatic protocol | 5-min supine rest; standing at ~1 and 3 min | Adopt pending UX/clinical review |

16. Clinical Evidence vs. Vitals Engineering Decisions

| Item | Classification | Rule |
| --- | --- | --- |
| Adult resting-rate reference / tachy-brady labels | Clinical reference | 60-100 BPM is a common adult resting reference; >100 and <60 are descriptive clinical terms with context-dependent exceptions. [2][3] |
| Marked-low <50 BPM reference | Clinical/context reference | Used only as a physician-facing reference; symptoms and clinical context determine significance. [4] |
| Orthostatic BP definition / posture protocol | Clinical reference | Use paired supine-standing HR/BP; OH reference definition comes from AHA statement. [5] |
| True HRV SDNN/RMSSD | Clinical/statistical reference | Requires validated beat-to-beat RR/NN intervals; do not compute from spot BPM. [9] |
| Heart-rate recovery | Research reference | Formula is valid only in a standardized exercise/recovery protocol; P2/deferred. [10] |
| Mean / SD / OLS / LOESS mathematics | Statistical reference | Shared deterministic implementations follow NIST methods. [6][7][8] |
| Trend/significance sample gates | Vitals engineering decision | Product stability gates, not diagnostic criteria. |
| Personal baseline windows | Vitals engineering decision | Recent 7 days vs preceding 30 days; configurable/versioned. |
| Symptom / medication association windows | Vitals engineering decision | Timestamp and pre/post windows support temporal review only; not causal inference. |
| 15-minute event clustering / 2-minute cross-vital fallback | Vitals engineering decision | Product grouping rules; preserve exact event timestamps and provenance. |

17. References

[1] Vitals Development Project. Mobile Application Development Roadmap - .NET MAUI | iOS & Android | September 2026 Launch Target. Internal project document, Sep 2026 v2.

[2] American Heart Association. All About Heart Rate. https://www.heart.org/en/health-topics/high-blood-pressure/the-facts-about-high-blood-pressure/all-about-heart-rate-pulse

[3] American Heart Association. Tachycardia: Fast Heart Rate; Bradycardia: Slow Heart Rate. Last reviewed Sep 24-25, 2024. https://www.heart.org/en/health-topics/arrhythmia/about-arrhythmia/tachycardia--fast-heart-rate ; https://www.heart.org/en/health-topics/arrhythmia/about-arrhythmia/bradycardia--slow-heart-rate

[4] Kusumoto FM, Schoenfeld MH, Barrett C, et al. 2018 ACC/AHA/HRS guideline on the evaluation and management of patients with bradycardia and cardiac conduction delay. Circulation. 2019;140:e382-e482. doi:10.1161/CIR.0000000000000628. https://www.ahajournals.org/doi/10.1161/CIR.0000000000000628

[5] Juraschek SP, Cortez MM, Flack JM, et al. Orthostatic Hypotension in Adults With Hypertension: A Scientific Statement From the American Heart Association. Hypertension. 2024;81:e16-e30. doi:10.1161/HYP.0000000000000236. https://www.ahajournals.org/doi/10.1161/HYP.0000000000000236

[6] NIST/SEMATECH e-Handbook of Statistical Methods. Linear Least Squares Regression / Least Squares. https://www.itl.nist.gov/div898/handbook/pmd/section1/pmd141.htm ; https://www.itl.nist.gov/div898/handbook/pmd/section4/pmd431.htm

[7] NIST/SEMATECH e-Handbook of Statistical Methods. 4.1.4.4 LOESS (aka LOWESS). https://www.itl.nist.gov/div898/handbook/pmd/section1/pmd144.htm

Clinical review checkpoint Before production release, confirm threshold wording, orthostatic presentation, symptom categories, and any patient-facing safety messages with a licensed clinician. The statistical formulas are deterministic engineering definitions; the interpretation and escalation language are clinical product decisions.

[8] NIST/SEMATECH e-Handbook of Statistical Methods. 1.3.5.6 Measures of Scale. https://www.itl.nist.gov/div898/handbook/eda/section3/eda356.htm

[9] Task Force of the European Society of Cardiology and the North American Society of Pacing and Electrophysiology. Heart rate variability: standards of measurement, physiological interpretation and clinical use. Circulation. 1996;93(5):1043-1065. doi:10.1161/01.CIR.93.5.1043. PMID:8598068. https://pubmed.ncbi.nlm.nih.gov/8598068/

[10] Cole CR, Blackstone EH, Pashkow FJ, Snader CE, Lauer MS. Heart-Rate Recovery Immediately after Exercise as a Predictor of Mortality. N Engl J Med. 1999;341:1351-1357. doi:10.1056/NEJM199910283411804. https://www.nejm.org/doi/full/10.1056/NEJM199910283411804

Reference maintenance: re-verify AHA/ACC/HRS guidance before release. If future guidelines change resting-rate or orthostatic definitions, preserve versioned reference profiles so historical analysis does not silently change.